CPB FMEA #10 roller pump failure to turn

The AmSECT Safety Committee

Contributor: Gary Grist  RN CCP, Kansas City, Missouri USA

Friends-

I have had quite a few experiences with roller pump failures, many of them with the arterial pump. Over the years my preparation to deal with roller pump failures has changed.  For example the first two pumps I ever used were consoles (built-in roller pumps). I didn’t like those because I had no free standing spare back-up pump that I could utilize when one pump broke down. And since the broken pump could not be easily replaced until after the case was completed, the entire unit was out of action until clinical engineering could get around to fixing it.  From then on I only purchased modular pumps so that I could buy free standing back-up roller pumps to replace any defective roller pump and still keep the entire unit operational.  As the years passed, the free standing roller pumps were no longer being built.  All the roller pumps were computer controlled and would only work with the power supply and computer connections to the main instrument stack.  In other words, I could no longer position a free standing roller pump on the floor next to a failing arterial roller pump, transfer the raceway to the back-up pump and turn it on using the 110v wall plug as power. Why not use one of the older, free-standing pumps, you ask?  We’ll I could have except for one thing.  Once the manufacturer stopped supporting those pumps, clinical engineering took them out of service.  Besides I always had a fear that using a free standing back-up roller pump that was not connected to the servo-regulators was just another accident waiting to happen.  Free standing, battery powered centrifugal pumps are still being manufactured today.  And since it is pretty hard to blow a circuit up with a C-pump, I decided to use them as my back-ups to a failed arterial roller pump. Secondary pumps are less critical and can usually be changed without significant jeopardy to the patient provided that assistance is available to help; like airline pilots – one to fly the plane the other to trouble shoot the problem.

I had an interesting discussion with one of the engineers who designed the pumps we were using some years back.  She was visiting our program at the invitation of our manufacturer’s representative to get our opinion on the new pumps we had recently purchased. I told her that there were several things I did not like about the pumps.  For one thing, changing settings and turning the alarms off and on required too many key strokes on the touch screens.  I told her during routine use, this was not a problem.  But during a crisis when things are happening fast and the perfusionist must move quickly, the numerous keys strokes could get cluttered and delay a proper response.  She said that none of the perfusionists interviewed during the design phase even mentioned that as a problem.  I asked her if she had ever thought that speed might be a critical factor in an emergency when operating the pump.  They did not feel that that would be important.

The other thing I mentioned go her was the difficulty in replacing the roller pumps quickly if one should fail.  She told me that should never be a consideration because these pumps were so reliable they would never break down at a critical moment. I thought to myself; “What was the last project you worked on – the Titanic?”  I informed her that we already had two pumps fail.  The touch screens on each pump had no protection on them and could easily be shattered by someone dropping a clamp or vial on them, which is what happened to ours.  She wondered out loud why anyone would ‘deliberately drop’ a clamp or vial on the screens.  I asked her if that scenario had ever been considered when the pumps were being designed.  It was not.  For my pumps I cut some clear, rigid plastic covers to fit like trap doors with tape for hinges over each touch screen. That solved the problem on our end.

Roller pumps are reliable.  But that is no excuse not to be prepared to change one out.  There should always be a back-up unit readily available in or near the OR with back-up personnel to help change it.

This is one of the few FMEAs I have worked on where having immediate assistance available does not reduce the Harmfulness RPN, which has a maximum of 5.  If a lone perfusionist with no help had to deal with this failure, I would make this Harmfulness RPN a 6, if that were an option.

-Gary Grist

This week’s Failure Mode is below:

- Failure Mode: Roller pump failure to turn.

II. Potential Effects of Failure:

EFFECT:

1. Failure to initiate CPB or unintentional termination of CPB if arterial roller pump fails.

a. No blood being delivered to patient

b. Hypotension

c. Acidosis

d. Hypercapnea

e. Hypoxia

f. Need to hand crank pump

g. Organ failure

h. Death

2. Failure to initiate cardioplegia, ultrafiltration, ventricular venting or field suckers if secondary pumps fail.

III. Potential Cause of Failure:

CAUSE:

1.	Human error due unfamiliarity with equipment or inadequate training.

2. Internal mechanical or electrical malfunction

a. Power cable loose, disconnected or power supply failure

b. Internal overload tripped due to over occlusion

c. Pump motor, drive belt, main bearing or speed control failure.

3. Total OR power failure.

IV. Interventions to Prevent or Negate the Failure:

PRE-EMPTIVE MANAGEMENT:

1. Routine table top scenario discussion or wet lab simulation of a roller pump failure should be regularly performed at least annually or whenever new personnel or unfamiliar equipment is added. The discussion/simulation should include any servo-regulated alarm scenario that could disable the roller pump.

2. Use detailed rather than generalized checklists to reduce failures. Confirm checklist item with action accompanied by conscious out loud verbal repetition before check-off.

3. A flash light or other battery powered portable lighting should always be available ON THE PUMP and checked for proper function should there be an overall loss of power in the OR. Emergency flood lights in the OR do not satisfy this need.

4. All pump instrument stacks should have an uninterruptable DC battery power source in case the AC power source fails. Test for proper battery function and capacity during priming by temporarily disconnecting the AC power source.

5. Confirm by checklist the secure placement of wall plug and proper operation of individual components during set-up and prime.

6. During priming, listen for any unusual noises that might indicate an internal malfunction.

7. Adjust occlusion of each pump head according to accepted practices during set-up and prime.

8. Observe ‘Load’ light of each pump for excessive current use indicating over occlusion or internal power overload.

9. All pumps should undergo a manufacturer's or qualified clinical engineering routine maintenance regularly that specifically includes a battery check or replacement.

10. Hand cranks should be available for each pump head.

11. A back-up, portable, battery powered centrifugal pump should be available to emergently take over from a malfunctioning arterial roller pump. The process of replacing a modular arterial roller pump with a back-up roller pump may require too much time to safely perform.

12. Back-up roller pumps should be available to replace ancillary roller pumps.

13. Secondary personnel (perfusion assistants or clinical perfusionists) should always be nearby to obtain back-up equipment and assist in emergency procedures.

MANAGEMENT:

1.	 Power loss can be to the entire heart-lung unit or be localized to individual components of the heart-lung unit.

2.	 Immediately clamp venous line to prevent patient exsanguination in an arterial pump failure.

3.	 Begin hand cranking in the correct direction at appropriate RPMs before removing the venous line clamp.

4.	 Monitor flow with an independent flow meter.

5.	 The perfusionist and secondary personnel should each perform designated scenario-practiced tasks.

6. Check for displacement of electrical plug from wall power or at main connection to pump or circuit breaker if entire instrument stack becomes powerless.

7. If over occlusion or internal overload is suspected temporarily stop pump and loosen occlusion.

8. If occlusion mechanism is defective replace pump with back-up immediately.

9. If main bearing is frozen hand cranking will not be possible.  Replace with back-up unit immediately.

10. If individual pump is powerless, turn pump switch off, reset circuit breaker and turn pump back on.

11. If pump is still inoperative, replace with back-up unit.

12. If the arterial pump fails, clamp arterial and venous lines and transfer raceway to a back-up centrifugal pump by cutting the raceway and attaching to the C-pump head.

13. If speed controller failure occurs, pump may not respond to speed control knob and even “runaway” at maximum RPMs.  Turn off power to pump immediately, begin hand cranking. Replace with back-up unit.

14. Post-traumatic stress disorder therapy should be available if needed for the perfusionist and/or other surgery team members, particularly if the patient experiences an adverse outcome.

Note: Specialized cardioplegia pumps may have no manual operation option should they fail unless specialized hand cranks have been supplied by the manufacturer.

V. Risk Priority Number (RPN): (select the number from each category that you feel best categorizes the risk).

A. Severity (Harmfulness) Rating Scale: how detrimental can the failure be:

1) Slight, 2) Low, 3) Moderate, 4) High, 5) Critical

(The problems that this failure causes are 5, critical.)

B. Occurrence Rating Scale: how frequently does the failure occur:

1) Remote, 2) Low, 3) Moderate, 4) Frequent, 5) Very High

(This is a rare problem. So occurrence should be 1, remote.)

C. Detection Rating Scale: how easily the potential failure can be detected before it occurs:

1) Very High, 2) High, 3) Moderate, 4) Low, 5) Uncertain

(This problem may not be detectable before it occurs. So the detection RPN should be 5, uncertain.)

D. Patient Frequency Scale:

1) Only a small number of patients would be susceptible to this failure, 2) Many patients but not all would be susceptible to this failure, 3) All patients would be susceptible to this failure.

(This could happen to any patient. So the Patient Frequency RPN should be a 3.)

Multiply A*B*C*D = RPN.  The higher the RPN the more dangerous the Failure Mode.

The lowest risk would be 1*1*1*1* = 1. The highest risk would be 5*5*5*3 = 375. RPNs allow the perfusionist to prioritize the risk. Resources should be used to reduce the RPNs of higher risk failures first, if possible.

(The total RPN for this failure is 5*1*5*3 = 75.)